## Example
Report (4): pH
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
[5 empty] | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  |  | Data |  |  | Average
|  | a |  |  | 4 | 4.05 | 4.05 | 4.0333
|  | b |  |  | 4.05 | 4 | 4 | 4.0167
|  | c |  |  | 4.05 | 4.1 | 4.1 | 4.0833
|  | Average | [5 empty] | 4.0444
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
[5 empty] | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  |  | Data |  |  | Average
|  | d |  |  | 4.5 | 4.4 | 4.45 | 4.45
|  | e |  |  | 4.45 | 4.45 | 4.4 | 4.4333
|  | f |  |  | 4.4 | 4.4 | 4.4 | 4.4
|  | Average | [5 empty] | 4.4278
(4) Operational blank sample
Data
|  | g |  |  | 5.95
|  | h |  |  | 5.86
|  | i |  |  | 5.71
|  | Median |  |  | 5.86
3. Measurement conditions
|  |  |  | pH meter | Manufacturer: | AAAAAAAA
[5 empty] | Model: | BB-BB
[5 empty] | Years in use: | 8 | year(s)
|  |  |  | Electrode | Manufacturer: | AAAAAAAA
[5 empty] | Model: | DDDDD
[5 empty] | Years in use: | 1 | year(s)
[5 empty] | Display type: | 2 | ( Analogue: 1, Digital: 2 )
[5 empty] | Type of pH meter: | 1 | ( Desk top: 1, Potable: 2 )
[5 empty] | Water bath: | 1 | ( Used: 1, Not used: 0 )
[5 empty] | Temperature conversion to 25oC: | 1 | ( Yes: 1, No: 0 )
4. Calibration for pH measurement
|  |  |  | Standard solution |  | pH at 25oC | Chemicals
[5 empty] | pH 4: | 4.01 | C6H4(COOK)(COOH)
[5 empty] | pH 7: | 6.86 | KH2PO4,Na2HPO4
[5 empty] | Water temperature at the calibration: | 25 | oC
[5 empty] | Electric conductivity of de-ionized water: | 0.11 | mS m-1
Report (5): EC
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
[5 empty] | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  |  | Data [mS m-1] |  |  | Average
|  | a |  |  | 7.2 | 7.25 | 7.3 | 7.25
|  | b |  |  | 7.3 | 7.2 | 7.3 | 7.2667
|  | c |  |  | 7.3 | 7.25 | 7.2 | 7.25
|  | Average | [5 empty] | 7.2556
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
[5 empty] | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  |  | Data [mS m-1] |  |  | Average
|  | d |  |  | 2.9 | 3 | 3 | 2.9667
|  | e |  |  | 3.1 | 3 | 3 | 3.0333
|  | f |  |  | 3 | 3.1 | 3.1 | 3.0667
|  | Average | [5 empty] | 3.0222
(4) Operational blank sample
Data [mS m-1]
|  | g |  |  | 0.09
|  | h |  |  | 0.09
|  | i |  |  | 0.07
|  | Median |  |  | 0.09
3. Measurement conditions
|  |  |  | EC meter | Manufacturer: | AAAAAAAA
[5 empty] | Model: | BB-BB
[5 empty] | Years in use: | 8 | year(s)
|  |  |  | Cell | Manufacturer: | AAAAAAAA
[5 empty] | Model: | DDDDD
[5 empty] | Years in use: | 1 | year(s)
[5 empty] | Display type: | 2 | ( Analogue: 1, Digital: 2 )
[5 empty] | Type of EC meter: | 1 | ( Desk top: 1, Potable: 2 )
[5 empty] | Water bath: | 1 | ( Used: 1, Not used: 0 )
[5 empty] | Temperature conversion to 25oC: | 1 | ( Yes: 1, No: 0 )
4. Calibration for EC measurement
|  | Original standard solution |  |  | Chemical: | KCl
[5 empty] | Concentration: | 0.01 | mol L-1
Diluted solution for calibration
[6 empty] | Solution 1 | Solution 2 | Solution 3
|  |  |  | Concentration [mol/L] |  | 0.001 | 0.0005 | 0.0001
|  |  |  | Conductivity [mS/m] |  | 14.7 | 7.39 | 1.49
[5 empty] | Water temperature at the calibration: | 25 | oC
[5 empty] | Cell constant: | 0.976 | cm-1
Report (6): Sulfate ion (SO42-)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 80.8 | 86.5 | 78.3 | 81.8667
|  | b |  |  | 80.9 | 78.6 | 80.3 | 79.9333
|  | c |  |  | 81.1 | 81.5 | 81.7 | 81.4333
|  | Average | [5 empty] | 81.0778
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 25 | 25.4 | 25.1 | 25.1667
|  | e |  |  | 24.8 | 24.7 | 24 | 24.5
|  | f |  |  | 24 | 24.3 | 24.5 | 24.2667
|  | Average | [5 empty] | 24.6444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0.1
|  | i |  |  | 0
Median*
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry: 2, Others: 9)
|  |  |  | Manufacturer: | DDDDDDDDD
|  |  |  | Model: | DD-EEE
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | CC-AA |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2020 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | CC-GG |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2020
|  |  | Eluent: | 2.7mmol/l Na2CO3/ 0.3mmol NaHCO3
|  |  | Flow rate: | 1.2 | mL min-1
|  |  | Injection method: | 1 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 25 | μL
|  |  | Column temp. | 40 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
Suppressor: |  |  |  | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | EEEEEEEEEE
--->Year of installation: |  |  |  | 2020
4. Standard solution
|  |  | Chemicals: | Na2SO4 (SO42-)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)
Report (7): Nitrate ion (NO3-)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 93.6 | 93.4 | 93 | 93.3333
|  | b |  |  | 93.6 | 97.3 | 93.6 | 94.8333
|  | c |  |  | 93.6 | 93.4 | 95 | 94
|  | Average | [5 empty] | 94.0556
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 33.8 | 32.1 | 34.6 | 33.5
|  | e |  |  | 34.2 | 33.4 | 31.4 | 33
|  | f |  |  | 32.9 | 34.1 | 30.7 | 32.5667
|  | Average | [5 empty] | 33.0222
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.2
|  | h |  |  | 0
|  | I |  |  | 0.1
Median*
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry: 2, Others: 9)
|  |  |  | Manufacturer: | DDDDDDDDD
|  |  |  | Model: | DD-EEE
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | CC-AA |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2020 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | CC-GG |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2020
|  |  | Eluent: | 2.7mmol/l Na2CO3/ 0.3mmol NaHCO3
|  |  | Flow rate: | 1.2 | mL min-1
|  |  | Injection method: | 1 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 25 | μL
|  |  | Column temp. | 40 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | EEEEEEEEEE
--->Year of installation: |  |  |  | 2020
4. Standard solution
Chemicals:
Concentration:
|  |  | Unit of concentration: |  | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: |  | (Commercial: 1, Prepared at your laboratory: 2)
Report (8): Chloride ion (Cl-)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 99 | 99.5 | 99.8 | 99.4333
|  | b |  |  | 98.8 | 98.1 | 99.4 | 98.7667
|  | c |  |  | 99.3 | 99.1 | 99 | 99.1333
|  | Average | [5 empty] | 99.1111
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 40.1 | 40.8 | 38.5 | 39.8
|  | e |  |  | 38.9 | 40.5 | 40 | 39.8
|  | f |  |  | 38.4 | 38.6 | 39.2 | 38.7333
|  | Average | [5 empty] | 39.4444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.1
|  | h |  |  | 0.1
|  | i |  |  | 0.3
Median*
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry: 2, Others: 9)
|  |  |  | Manufacturer: | DDDDDDDDD
|  |  |  | Model: | DD-EEE
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | CC-AA |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2020 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | CC-GG |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2020
|  |  | Eluent: | 2.7mmol/l Na2CO3/ 0.3mmol NaHCO3
|  |  | Flow rate: | 1.2 | mL min-1
|  |  | Injection method: | 1 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 25 | μL
|  |  | Column temp. | 40 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | EEEEEEEEEE
--->Year of installation: |  |  |  | 2020
4. Standard solution
|  |  | Chemicals: | NaCl (Cl-)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)
Report (9): Ammonium ion (NH4+)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 75.9 | 78.4 | 78.7 | 77.6667
|  | b |  |  | 78.6 | 78.9 | 78.7 | 78.7333
|  | c |  |  | 78.4 | 78.4 | 78.7 | 78.5
|  | Average | [5 empty] | 78.3
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 25.3 | 24.1 | 24.7 | 24.7
|  | e |  |  | 24.8 | 28.9 | 24.7 | 26.1333
|  | f |  |  | 27.6 | 26.7 | 24.8 | 26.3667
|  | Average | [5 empty] | 25.7333
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.3
|  | h |  |  | 0.3
|  | i |  |  | 0.2
|  | Median* |  |  | 0.3
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry[Indophenol]: 2, Spectrophotometry[Other]: 3, Others: 9)
|  |  |  | Manufacturer: | DDDDDDDDD
|  |  |  | Model: | DD-EEE
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | CCCCC |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2020 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | DDDDD |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2020
|  |  | Eluent: | 20mmol/l CH3SO3H
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 1 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 25 | μL
|  |  | Column temp. | 40 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | FFFFFF
--->Year of installation: |  |  |  | 2020
4. Standard solution
|  |  | Chemicals: | NH4(NO3) (NH4+)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)
Report (10): Sodium ion (Na+)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 80.5 | 76.9 | 78.4 | 78.6
|  | b |  |  | 79.6 | 78.2 | 79.5 | 79.1
|  | c |  |  | 78.8 | 80.4 | 76.3 | 78.5
|  | Average | [5 empty] | 78.7333
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 20.3 | 24.6 | 25.4 | 23.4333
|  | e |  |  | 25.7 | 25.6 | 24.7 | 25.3333
|  | f |  |  | 24.3 | 25.4 | 27.6 | 25.7667
|  | Average | [5 empty] | 24.8444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.7
|  | h |  |  | 0.5
|  | i |  |  | 0.8
Median*
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | DDDDDDDDD
|  |  |  | Model: | DD-EEE
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: |  |  |  | CCCCC |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2020 |  | Correction: |  | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: | DDDDD |  | Cs solution: |  | (Added: 1, No: 0)
--->Year of installation: |  |  |  | 2020 |  | Flame:
|  |  | Eluent: | 20mmol/l CH3SO3H |  |  |  |  | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 1 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 25 | μL
|  |  | Column temp. | 40 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | FFFFFF
--->Year of installation: |  |  |  | 2020
4. Standard solution
|  |  | Chemicals: | NaCl (Na+)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)
Report (11): Potassium ion (K+)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 12 | 11.5 | 12.5 | 12
|  | b |  |  | 13.6 | 12.9 | 12.7 | 13.0667
|  | c |  |  | 12.4 | 12.7 | 12.4 | 12.5
|  | Average | [5 empty] | 12.5222
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 10.2 | 10.3 | 10.2 | 10.2333
|  | e |  |  | 10.3 | 10.1 | 9.25 | 9.9
|  | f |  |  | 9.62 | 10.8 | 10.1 | 10.1667
|  | Average | [5 empty] | 10.1
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.2
|  | h |  |  | 0.2
|  | i |  |  | 0.2
Median*
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | DDDDDDDDD
|  |  |  | Model: | DD-EEE
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: |  |  |  | CCCCC |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2020 |  | Correction: |  | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: | DDDDD |  | Cs solution: |  | (Added: 1, No: 0)
--->Year of installation: |  |  |  | 2020 |  | Flame:
|  |  | Eluent: | 20mmol/l CH3SO3H |  |  |  |  | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 1 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 25 | μL
|  |  | Column temp. | 40 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | FFFFFF
--->Year of installation: |  |  |  | 2020
4. Standard solution
|  |  | Chemicals: | KCl (K+)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)
Report (12): Calcium ion (Ca2+)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 32 | 31.4 | 35.1 | 32.8333
|  | b |  |  | 34.9 | 38.6 | 38.8 | 37.4333
|  | c |  |  | 36.3 | 35.2 | 35.6 | 35.7
|  | Average | [5 empty] | 35.3222
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 12.3 | 12.4 | 12.6 | 12.4333
|  | e |  |  | 12.5 | 12.5 | 12.3 | 12.4333
|  | f |  |  | 12.4 | 12.6 | 12.4 | 12.4667
|  | Average | [5 empty] | 12.4444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.8
|  | h |  |  | 1.2
|  | i |  |  | 1
Median*
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 2 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | HHHHHHHHHHHHHHHHHH
|  |  |  | Model: | IIIIIIIIII
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: | [5 empty] | Wave length: | 423 | nm
--->Year of installation: | [5 empty] | Correction: | 1 | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: |  |  | La solution: | 1 | (Added: 1, No: 0)
--->Year of installation: | [5 empty] | Flame: | 1
|  |  | Eluent: | [5 empty] | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: |  | mL min-1
|  |  | Injection method: |  | (Auto: 1, Manual: 2)
|  |  | Injection volume: |  | μL
|  |  | Column temp. |  | oC
Kind of detector: |  |  |  |  | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: |  | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | CaCO3 (Ca2+)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)
Report (13): Magnesium ion (Mg2+)
1. Person in change of measurement
| Name: | Dr./Mr./Ms. A B
[6 empty] | Years of experience for chemical analysis: | 10 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-01 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-07 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 13.6 | 15.4 | 15.2 | 14.7333
|  | b |  |  | 13.9 | 15.1 | 14.8 | 14.6
|  | c |  |  | 14.7 | 14.5 | 14.2 | 14.4667
|  | Average | [5 empty] | 14.6
| (3) No.212w |  |  |  | Serial No. | 122 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 3.51 | 3.1 | 3.5 | 3.3667
|  | e |  |  | 3 | 2.9 | 3.11 | 3
|  | f |  |  | 3.1 | 3.02 | 3.4 | 3.1667
|  | Average | [5 empty] | 3.1778
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0.1
|  | h |  |  | 0.1
|  | i |  |  | 0.2
|  | Median* |  |  | 0.1
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 2 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | HHHHHHHHHHHHHHHHHH
|  |  |  | Model: | IIIIIIIIII
|  |  |  | Year of manufacture: | 2019
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: | [5 empty] | Wave length: | 285 | nm
--->Year of installation: | [5 empty] | Correction: | 1 | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: |  |  | La solution: | 1 | (Added: 1, No: 0)
--->Year of installation: | [5 empty] | Flame: | 1
|  |  | Eluent: | [5 empty] | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: |  | mL min-1
|  |  | Injection method: |  | (Auto: 1, Manual: 2)
|  |  | Injection volume: |  | μL
|  |  | Column temp. |  | oC
Kind of detector: |  |  |  |  | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: |  | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | Mg(NO3)2 (Mg2+)
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (1)
Report (1): Summary
1. Organization / Main Analyst
|  | Country: | Vietnam
|  | Organization: | Environmental Laboratory - Center for Environmental Research - IMHEN- MoNRE
|  | Main Analyst: | Nguyen Truong Giang
|  | TEL: | (84.04).37733090-509
|  | E-mail: | nguyenkimanh1004@gmail.com
2. Summary of results
The calculation of R1 and R2 is based on the manual and QA/QC program. If the value of "R1" or "R2" does not meet the required criteria, please confirm their analytical results again.
| No. 211w | pH | EC | SO42- | NO3- | Cl- | NH4+ | Na+ | K+ | Ca2+ | Mg2+ | R1 | R2
|  | - | mS m-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | Ion balance | EC check
| a | 5.0467 | 2.9167 | 49.6333 | 39.0667 | 43.3333 | 55.3667 | 34.2333 | 6.5333 | 23.5667 | 11.1 | -2.0271 | -3.0803
| b | 5.0533 | 2.9067 | 49.5333 | 38.6667 | 43.4 | 54.9667 | 34.2 | 6.5667 | 23.4667 | 10.8 | -2.2646 | -3.2004
| c | 5.0567 | 2.91 | 49.3667 | 38.7667 | 43.1667 | 55.5 | 34.4 | 6.4 | 23.5 | 10.8667 | -1.9343 | -3.2814
| Average | 5.0522 | 2.9111 | 49.5111 | 38.8333 | 43.3 | 55.2778 | 34.2778 | 6.5 | 23.5111 | 10.9222 | -2.0753 | -3.1875
| Blank sample | 6.25 | 0.01 | 0.028853 | 0.038243 | 0.033541 | 0.038827 | 0.0502 | 0.031105 | 0.069875 | 0.057511
| No.212w | pH | EC | SO42- | NO3- | Cl- | NH4+ | Na+ | K+ | Ca2+ | Mg2+ | R1 | R2
|  | - | mS m-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | Ion balance | EC check
| d | 5.49 | 1.1867 | 19.9667 | 13.2333 | 18.8667 | 23.6333 | 14.1 | 3.9667 | 8.6667 | 4 | -1.2397 | -4.2414
| e | 5.4867 | 1.2 | 19.9333 | 13.3 | 18.5 | 23.3667 | 14.2 | 4.0667 | 8.5333 | 3.8667 | -1.3953 | -5.0607
| f | 5.4833 | 1.1933 | 19.7667 | 13.3333 | 18.2667 | 23.4 | 14.1333 | 4.0333 | 9.1 | 3.9667 | -0.10367 | -4.5928
| Average | 5.4867 | 1.1933 | 19.8889 | 13.2889 | 18.5444 | 23.4667 | 14.1444 | 4.0222 | 8.7667 | 3.9444 | -0.91222 | -4.6322
| Blank sample | 6.25 | 0.01 | 0.028853 | 0.038243 | 0.033541 | 0.038827 | 0.0502 | 0.031105 | 0.069875 | 0.057511

## (2)
Report (2): Concentration of standard solution
1. Organization
| Country: | Vietnam
| Organization: | Environmental Laboratory - Center for Environmental Research - IMHEN- MoNRE
2. Concentrations of ions in the standard solutions
Please fill the prepared concentrations of ions in standard solutions used for drawing calibration curves in the table below.
The concentration values shall be rounded to the 2nd decimal place.
In addition, please input "1" in each cell of the standard solution which used for the analysis of sample No. 211w and No.212w, respectively.
| Standard solution | SO42- | No. 211w | No.212w | NO3- | No. 211w | No.212w | Cl- | No. 211w | No.212w | NH4+ | No. 211w | No.212w | Na+ | No. 211w | No.212w | K+ | No. 211w | No.212w | Ca2+ | No. 211w | No.212w | Mg2+ | No. 211w | No.212w
|  | μmol L-1 |  |  | μmol L-1 |  |  | μmol L-1 |  |  | μmol L-1 |  |  | μmol L-1 |  |  | μmol L-1 |  |  | μmol L-1 |  |  | μmol L-1
| Std0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1
| Std1 | 1.04 | 1 | 1 | 0.805 | 1 | 1 | 1.41 | 1 | 1 | 2.77 | 1 | 1 | 2.175 | 1 | 1 | 1.28 | 1 | 1 | 2.495 | 1 | 1 | 2.055 | 1 | 1
| Std2 | 2.08 | 1 | 1 | 1.6126 | 1 | 1 | 2.8209 | 1 | 1 | 5.5432 | 1 | 1 | 4.3497 | 1 | 1 | 2.5575 | 1 | 1 | 4.99 | 1 | 1 | 4.1118 | 1 | 1
| Std3 | 4.16 | 1 | 1 | 3.2253 | 1 | 1 | 5.6417 | 1 | 1 | 11.0865 | 1 | 1 | 8.6994 | 1 | 1 | 5.1151 | 1 | 1 | 9.98 | 1 | 1 | 8.2237 | 1 | 1
| Std4 | 10.41 | 1 | 1 | 8.0632 | 1 | 1 | 14.1044 | 1 | 1 | 27.7162 | 1 | 1 | 21.7486 | 1 | 1 | 12.7877 | 1 | 1 | 24.9501 | 1 | 1 | 20.5592 | 1 | 1
| Std5 | 20.82 | 1 | 1 | 16.1264 | 1 | 1 | 28.2087 | 1 | 1 | 55.4324 | 1 | 1 | 43.4972 | 1 | 1 | 25.5754 | 1 | 1 | 49.9002 | 1 | 1 | 41.1184 | 1 | 1
| Std6 | 41.64 | 1 | 1 | 32.2529 | 1 | 1 | 56.4175 | 1 | 1 | 110.8647 | 1 | 1 | 86.9943 | 1 | 1 | 51.1509 | 1 | 1 | 99.8004 | 1 | 1 | 82.2368 | 1 | 1
| Std7 | 104.1 | 1 | 1 | 80.6322 | 1 | 1 | 141.0437 | 1 | 1 | 277.1619 | 1 | 1 | 217.4859 | 1 | 1 | 127.8772 | 1 | 1 | 249.501 | 1 | 1 | 205.5921 | 1 | 1
Std8
Std9
Std10

## (3)
Report (3): Detection limit and determination limit
1. Organization
| Country: | Vietnam
| Organization: | Environmental Laboratory - Center for Environmental Research - IMHEN- MoNRE
2. Not Detected (3s) and Lowest Determination Limit (10s)
Please fill each analytical value of Std1 obtained by repeating analysis 5 times in the blue cell.
Enter the values before rounding. If the number of significant digits exceeds 5, please round off the 6th significant digit to 5 significant digits.
If detection limits or determination limits do not meet their DQOs, please confirm your measurement conditions again.
| Analysis times | SO42- | NO3- | Cl- | NH4+ | Na+ | K+ | Ca2+ | Mg2+
|  | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1 | μmol L-1
| 1 | 1.06 | 0.78 | 1.36 | 2.79 | 2.15 | 1.26 | 2.52 | 2.09
| 2 | 1.08 | 0.82 | 1.38 | 2.8 | 2.21 | 1.3 | 2.45 | 2.12
| 3 | 1.03 | 0.83 | 1.42 | 2.75 | 2.2 | 1.29 | 2.49 | 2.05
| 4 | 1.05 | 0.8 | 1.4 | 2.74 | 2.17 | 1.25 | 2.53 | 2.08
| 5 | 1.07 | 0.77 | 1.39 | 2.76 | 2.13 | 1.28 | 2.42 | 2.15
| Average | 1.058 | 0.8 | 1.39 | 2.768 | 2.172 | 1.276 | 2.482 | 2.098
| Standard deviation (s.d.) | 0.019235 | 0.025495 | 0.022361 | 0.025884 | 0.033466 | 0.020736 | 0.046583 | 0.038341
| Detection Limit (3 s.d.) | 0.057706 | 0.076485 | 0.067082 | 0.077653 | 0.1004 | 0.062209 | 0.13975 | 0.11502
| Determination Limit (10 s.d.) | 0.19235 | 0.25495 | 0.22361 | 0.25884 | 0.33466 | 0.20736 | 0.46583 | 0.38341
| Concentration values of Std1 | 1.04 | 0.805 | 1.41 | 2.77 | 2.175 | 1.28 | 2.495 | 2.055

## (4)
Report (4): pH
1. Person in change of measurement
| Name: | Vu Xuan Hung
[5 empty] | Years of experience for chemical analysis: | 5 | year(s)
[5 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
|  |  |  | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  | Serial No. | 75 | : Printed on the label of bottle
|  |  |  | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  | Data |  |  | Average
|  | a |  | 5.05 | 5.05 | 5.04 | 5.0467
|  | b |  | 5.06 | 5.04 | 5.06 | 5.0533
|  | c |  | 5.06 | 5.05 | 5.06 | 5.0567
|  | Average |  |  |  |  | 5.0522
| (3) No.212w |  |  | Serial No. | 75 | : Printed on the label of bottle
|  |  |  | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  | Data |  |  | Average
|  | d |  | 5.48 | 5.49 | 5.5 | 5.49
|  | e |  | 5.5 | 5.47 | 5.49 | 5.4867
|  | f |  | 5.47 | 5.5 | 5.48 | 5.4833
|  | Average |  |  |  |  | 5.4867
(4) Operational blank sample
Data
|  | g |  | 6.26
|  | h |  | 6.25
|  | i |  | 6.23
|  | Median |  | 6.25
3. Measurement conditions
|  |  | pH meter | Manufacturer: | Metrohm
|  |  |  | Model: | 781
|  |  |  | Years in use: | 18 | year(s)
|  |  | Electrode | Manufacturer: | Metrohm
|  |  |  | Model: | Metrohm 6.0.258.010
|  |  |  | Years in use: | 9 | year(s)
|  |  |  | Display type: | 2 | ( Analogue: 1, Digital: 2 )
|  |  |  | Type of pH meter: | 1 | ( Desk top: 1, Potable: 2 )
|  |  |  | Water bath: | 1 | ( Used: 1, Not used: 0 )
|  |  |  | Temperature conversion to 25oC: | 1 | ( Yes: 1, No: 0 )
4. Calibration for pH measurement
|  |  | Standard solution |  | pH at 25oC | Chemicals
|  |  |  | pH 4: | 4.01 | Merck
|  |  |  | pH 7: | 6.98 | Merck
|  |  |  | Water temperature at the calibration: | 25 | oC
|  |  |  | Electric conductivity of de-ionized water: | 0.01 | mS m-1

## (5)
Report (5): EC
1. Person in change of measurement
| Name: | Vu Xuan Hung
[5 empty] | Years of experience for chemical analysis: | 5 | year(s)
[5 empty] | Years of experience for chemical analysis of acid deposition: | 5 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
|  |  |  | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  | Serial No. | 75 | : Printed on the label of bottle
|  |  |  | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  | Data [mS m-1] |  |  | Average
|  | a |  | 2.93 | 2.9 | 2.92 | 2.9167
|  | b |  | 2.91 | 2.89 | 2.92 | 2.9067
|  | c |  | 2.93 | 2.88 | 2.92 | 2.91
|  | Average |  |  |  |  | 2.9111
| (3) No.212w |  |  | Serial No. | 75 | : Printed on the label of bottle
|  |  |  | Water temperature at measurement: | 25 | oC
|  | Diluted sample |  | Data [mS m-1] |  |  | Average
|  | d |  | 1.21 | 1.17 | 1.18 | 1.1867
|  | e |  | 1.22 | 1.18 | 1.2 | 1.2
|  | f |  | 1.2 | 1.19 | 1.19 | 1.1933
|  | Average |  |  |  |  | 1.1933
(4) Operational blank sample
Data [mS m-1]
|  | g |  | 0.01
|  | h |  | 0.01
|  | i |  | 0.01
|  | Median |  | 0.01
3. Measurement conditions
|  |  | EC meter | Manufacturer: | YSI
|  |  |  | Model: | 3200
|  |  |  | Years in use: | 18 | year(s)
|  |  | Cell | Manufacturer: | YSI
|  |  |  | Model: | YSI 3252
|  |  |  | Years in use: | 18 | year(s)
|  |  |  | Display type: | 1 | ( Analogue: 1, Digital: 2 )
|  |  |  | Type of EC meter: | 1 | ( Desk top: 1, Potable: 2 )
|  |  |  | Water bath: | 1 | ( Used: 1, Not used: 0 )
|  |  |  | Temperature conversion to 25oC: | 1 | ( Yes: 1, No: 0 )
4. Calibration for EC measurement
|  | Original standard solution |  | Chemical: | HANNA
|  |  |  | Concentration: |  | mol L-1
Diluted solution for calibration
[5 empty] | Solution 1 | Solution 2 | Solution 3
Concentration [mol/L]
|  |  | Conductivity [mS/m] |  | 8.4
|  |  |  | Water temperature at the calibration: | 25 | oC
|  |  |  | Cell constant: | 1 | cm-1

## (6)
Report (6): Sulfate ion (SO42-)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 49.8 | 49.4 | 49.7 | 49.6333
|  | b |  |  | 50.1 | 48.9 | 49.6 | 49.5333
|  | c |  |  | 50.5 | 48.4 | 49.2 | 49.3667
|  | Average | [5 empty] | 49.5111
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 20.2 | 20 | 19.7 | 19.9667
|  | e |  |  | 19.8 | 20.1 | 19.9 | 19.9333
|  | f |  |  | 20.2 | 19.6 | 19.5 | 19.7667
|  | Average | [5 empty] | 19.8889
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.028853
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry: 2, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | Metrosep - Anion dual 2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2018 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | Metrosep RP2guard/3.5 |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2019
|  |  | Eluent: | NaHCO3 &Na2CO3
|  |  | Flow rate: | 0.8 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
Suppressor: |  |  |  | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | 753 Suppressor Module - Metrohm
--->Year of installation: |  |  |  | 2000
4. Standard solution
|  |  | Chemicals: | Na2SO4
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (7)
Report (7): Nitrate ion (NO3-)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 39.3 | 38.9 | 39 | 39.0667
|  | b |  |  | 38.6 | 38.5 | 38.9 | 38.6667
|  | c |  |  | 38.7 | 38.4 | 39.2 | 38.7667
|  | Average | [5 empty] | 38.8333
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 13.4 | 12.9 | 13.4 | 13.2333
|  | e |  |  | 13.3 | 13.1 | 13.5 | 13.3
|  | f |  |  | 13.1 | 13.4 | 13.5 | 13.3333
|  | Average | [5 empty] | 13.2889
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | I |  |  | 0
|  | Median* |  |  | 0.038243
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry: 2, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | Metrosep - Anion dual 2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2018 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | Metrosep RP2guard/3.5 |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2019
|  |  | Eluent: | NaHCO3 &Na2CO3
|  |  | Flow rate: | 0.8 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | 753 Suppressor Module - Metrohm
--->Year of installation: |  |  |  | 2001
4. Standard solution
|  |  | Chemicals: | NaNO3
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (8)
Report (8): Chloride ion (Cl-)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 43.5 | 43.1 | 43.4 | 43.3333
|  | b |  |  | 42.9 | 43.6 | 43.7 | 43.4
|  | c |  |  | 43 | 43.4 | 43.1 | 43.1667
|  | Average | [5 empty] | 43.3
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 19.2 | 18.4 | 19 | 18.8667
|  | e |  |  | 18.5 | 18.8 | 18.2 | 18.5
|  | f |  |  | 18.1 | 18.3 | 18.4 | 18.2667
|  | Average | [5 empty] | 18.5444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.033541
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry: 2, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | Metrosep - Anion dual 2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2018 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | Metrosep RP2guard/3.5 |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2019
|  |  | Eluent: | NaHCO3 &Na2CO3
|  |  | Flow rate: | 0.8 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 1 | (Attached: 1, No: 0)
--->Model of suppressor: |  |  |  | 53 Suppressor Module - Metrohm
--->Year of installation: |  |  |  | 2001
4. Standard solution
|  |  | Chemicals: | NaCl
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (9)
Report (9): Ammonium ion (NH4+)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 55.6 | 55.1 | 55.4 | 55.3667
|  | b |  |  | 54.9 | 55.4 | 54.6 | 54.9667
|  | c |  |  | 55.8 | 55.5 | 55.2 | 55.5
|  | Average | [5 empty] | 55.2778
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 23.7 | 23.4 | 23.8 | 23.6333
|  | e |  |  | 22.9 | 23.5 | 23.7 | 23.3667
|  | f |  |  | 23.1 | 23.4 | 23.7 | 23.4
|  | Average | [5 empty] | 23.4667
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.038827
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion chromatography: 1, Spectrophotometry[Indophenol]: 2, Spectrophotometry[Other]: 3, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Spectrophotometry >
Model of Analytical column: |  |  |  | Metrosep cation 1-2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2019 |  | Heating temp. |  | oC
Model of Guard column: |  |  | Guard column: | Metrosep C6 S- guard/2.0 |  | Heating time: |  | hour(s)
--->Year of installation: |  |  |  | 2019
|  |  | Eluent: | Acid Tactric & Acid Dipinilic
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 0 | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | NH4Cl
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (10)
Report (10): Sodium ion (Na+)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 33.9 | 34.5 | 34.3 | 34.2333
|  | b |  |  | 33.7 | 34.8 | 34.1 | 34.2
|  | c |  |  | 33.6 | 35.1 | 34.5 | 34.4
|  | Average | [5 empty] | 34.2778
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 14.3 | 13.9 | 14.1 | 14.1
|  | e |  |  | 13.8 | 14.5 | 14.3 | 14.2
|  | f |  |  | 14.1 | 14.4 | 13.9 | 14.1333
|  | Average | [5 empty] | 14.1444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.0502
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: |  |  |  | Metrosep cation 1-2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2019 |  | Correction: |  | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: | Metrosep C6 S- guard/2.0 |  | Cs solution: |  | (Added: 1, No: 0)
--->Year of installation: |  |  |  | 2019 |  | Flame:
|  |  | Eluent: | Acid Tactric & Acid Dipinilic |  |  |  |  | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 0 | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | NaCl
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (11)
Report (11): Potassium ion (K+)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 6.8 | 6.5 | 6.3 | 6.5333
|  | b |  |  | 7.1 | 6.2 | 6.4 | 6.5667
|  | c |  |  | 6.5 | 6.3 | 6.4 | 6.4
|  | Average | [5 empty] | 6.5
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 3.9 | 4.2 | 3.8 | 3.9667
|  | e |  |  | 4.1 | 4.4 | 3.7 | 4.0667
|  | f |  |  | 3.8 | 3.7 | 4.6 | 4.0333
|  | Average | [5 empty] | 4.0222
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.031105
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: |  |  |  | Metrosep cation 1-2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2019 |  | Correction: |  | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: | Metrosep C6 S- guard/2.0 |  | Cs solution: |  | (Added: 1, No: 0)
--->Year of installation: |  |  |  | 2019 |  | Flame:
|  |  | Eluent: | Acid Tactric & Acid Dipinilic |  |  |  |  | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 0 | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | KCl
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (12)
Report (12): Calcium ion (Ca2+)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 23.8 | 23.3 | 23.6 | 23.5667
|  | b |  |  | 24.1 | 23.4 | 22.9 | 23.4667
|  | c |  |  | 23.6 | 23.5 | 23.4 | 23.5
|  | Average | [5 empty] | 23.5111
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 8.5 | 8.9 | 8.6 | 8.6667
|  | e |  |  | 8.4 | 8.2 | 9 | 8.5333
|  | f |  |  | 9.1 | 9.4 | 8.8 | 9.1
|  | Average | [5 empty] | 8.7667
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.069875
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: |  |  |  | Metrosep cation 1-2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2019 |  | Correction: |  | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: | Metrosep C6 S- guard/2.0 |  | La solution: |  | (Added: 1, No: 0)
--->Year of installation: |  |  |  | 2019 |  | Flame:
|  |  | Eluent: | Acid Tactric & Acid Dipinilic |  |  |  |  | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 0 | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | CaCl2
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (13)
Report (13): Magnesium ion (Mg2+)
1. Person in change of measurement
| Name: | Nguyen Truong Giang
[6 empty] | Years of experience for chemical analysis: | 8 | year(s)
[6 empty] | Years of experience for chemical analysis of acid deposition: | 8 | year(s)
2. Result of measurement
| (1) Date of measurement etc. |  |  |  | Start date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
[5 empty] | Finish date: | 2021-12-02 | : Year/Month/Day (ex. 2020/10/31)
| (2) No. 211w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | a |  |  | 11.3 | 11.1 | 10.9 | 11.1
|  | b |  |  | 10.5 | 10.9 | 11 | 10.8
|  | c |  |  | 10.9 | 10.7 | 11 | 10.8667
|  | Average | [5 empty] | 10.9222
| (3) No.212w |  |  |  | Serial No. | 75 | : Printed on the label of bottle
|  | Diluted sample |  |  | Data [μmol L-1] |  |  | Average
|  | d |  |  | 4.3 | 4 | 3.7 | 4
|  | e |  |  | 3.8 | 4.2 | 3.6 | 3.8667
|  | f |  |  | 3.9 | 4.2 | 3.8 | 3.9667
|  | Average | [5 empty] | 3.9444
(4) Operational blank sample
Data[μmol L-1]
|  | g |  |  | 0
|  | h |  |  | 0
|  | i |  |  | 0
|  | Median* |  |  | 0.057511
* In the case that the median is less than the detection limit, 1/2 detection limit is automatically adopted as the result of the operational blank sample.
3. Measurement conditions
|  |  | Analytical method: | 1 |  | If "9", enter the method used >>>
(Ion Chromatography: 1, Atomic Absorption Spectrometry: 2, Atomic Emission Spectrometry: 3, Others: 9)
|  |  |  | Manufacturer: | Metrohm
|  |  |  | Model: | Mic 3
|  |  |  | Year of manufacture: | 1998
< Ion Chromatography > | [5 empty] | < Atomic Absorption/Emission Spectrometry >
Model of Analytical column: |  |  |  | Metrosep cation 1-2 |  | Wave length: |  | nm
--->Year of installation: |  |  |  | 2019 |  | Correction: |  | (No: 1, Zeeman: 2, D2: 3)
Model of Guard column: |  |  | Guard column: | Metrosep C6 S- guard/2.0 |  | La solution: |  | (Added: 1, No: 0)
--->Year of installation: |  |  |  | 2019 |  | Flame:
|  |  | Eluent: | Acid Tactric & Acid Dipinilic |  |  |  |  | (Air-C2H2: 1, N2O-C2H2: 2, Others: 9)
|  |  | Flow rate: | 1 | mL min-1
|  |  | Injection method: | 2 | (Auto: 1, Manual: 2)
|  |  | Injection volume: | 20 | μL
|  |  | Column temp. | 35 | oC
Kind of detector: |  |  |  | 1 | (EC: 1, Others: 9)
If others, enter the kind of detector:
|  |  | Suppressor: | 0 | (Attached: 1, No: 0)
--->Model of suppressor:
--->Year of installation:
4. Standard solution
|  |  | Chemicals: | MgCl2
|  |  | Concentration: | 1000
|  |  | Unit of concentration: | 1 | (mg L-1: 1, μmol L-1: 2)
|  |  | Preparation method: | 1 | (Commercial: 1, Prepared at your laboratory: 2)

## (14)
Report (14): Information on QA/QC in the laboratory
1. Utilization of Standard operating procedures (SOPs)
| 1 | (Use: 1, Not use: 0)
2. System of Data Checking
(1) National QA/QC Manager approve data submission
| 1 | (Yes: 1, No: 0)
(2) Main analyst in the laboratory approve data submission
| 1 | (Yes: 1, No: 0)
(3) Others approve data submission
| 0 | (Yes: 1, No: 0)
Enter others name
3. Quality control for Analysis on monitoring activities
(1) Participation for other Inter-laboratory comparison project
a. World Meteorological Organization Global Atmosphere Watch (WMO/GAW)
(Yes: 1, No: 0)
b. U.S. Geological Survey (USGS)
(Yes: 1, No: 0)
c. European Monitoring and Evaluation Programme (EMEP)
(Yes: 1, No: 0)
d. No participating
(Yes: 1, No: 0)
4. Operation in the laboratory
(1) Preparation of "Diluted ILC sample (211w, 212w)", in this attempt.
a. How long did it take to finish the measurement after dilution?
(a few days: 1, a week: 2, a month: 3, a few months: 4, Others: 5)
Enter others period
b. Material of container to store Diluted ILC sample
(Plastic: 1, Glass: 2, Others: 3)
Enter others name
c. Storage condition of Diluted ILC sample
(In refrigerators: 1, At room temperature: 2, Others: 3)
Enter others name
(2) Preparation of "Diluted standard solution (STD)"
a. Frequency of Preparation
(every time: 1, every week: 2, every month: 3, every year: 4, Others: 6)
Enter others frequency
b. Material of container to store Diluted STD
(Plastic: 1, Glass: 2, Others: 3)
Enter others name
c. Storage condition of Diluted STD
(In refrigerators: 1, At room temperature: 2, Others: 3)
Enter others name
d. Are these answers in "(2) a ～c" regulated in your SOPs?
(Yes: 1, No: 0, Partially: 3)
(3) Preparation of "Standard Reference Materials (RM)".
a. Frequency of Preparation
(every time: 1, every week: 2, every month: 3, every year: 4, Others: 6)
Enter others frequency
b. Material of container to store Diluted RM
(Plastic: 1, Glass: 2, Others: 3)
Enter others name
c. Storage condition of Diluted RM
(In refrigerators: 1, At room temperature: 2, Others: 3)
Enter others name
d. Are these answers in "(3) a ～c" regulated in your SOPs?
(Yes: 1, No: 0, Partially: 3)
(4) Preparation of "Eluent for IC" (If you don't use IC, please skip (4).)
a. Frequency of Preparation
(every time: 1, every week: 2, every month: 3, every year: 4, Others: 6)
Enter others frequency
b. Material of container to store prepared eluent
(Plastic: 1, Glass: 2, Others: 3)
Enter others name
c. Storage condition of prepared eluent
(In refrigerators: 1, At room temperature: 2, Others: 3)
Enter others name
d. Are these answers in "(4) a ～c" regulated in your SOPs?
(Yes: 1, No: 0, Partially: 3)
5. Problems on monitoring activities
If there are any problems on wet deposition monitoring in your laboratory, please write in the box below.

## Graph & Chart
< Please append the copy of calibration curve > | < Please append the copy of analysis charts of Std1 |  | < Please append one copy of analysis charts of the operational blank sample |  | < Please append one copy of analysis charts of the Sample No. 211w |  | < Please append one copy of analysis charts of the Sample No.212w
|  | for SO42-, NO3-, Cl-, NH4+, Na+, K+, Ca2+, Mg2+ > |  | for SO42-, NO3-, Cl-, NH4+, Na+, K+, Ca2+, Mg2+ > |  | for SO42-, NO3-, Cl-, NH4+, Na+, K+, Ca2+, Mg2+ > |  | for SO42-, NO3-, Cl-, NH4+, Na+, K+, Ca2+, Mg2+ >

## List (hidden)
| name | exp1 | exp2 | start_date | finish_date | pH of blank sample | h_serial_no | h_water_tmp | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_water_tmp | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | manfacturer_pHmeter | model_ pHmeter | years_pHmeter | manufacturer_electrode | model_electrode | years_electrode | flg_disp | type_disp | flg_meter | type_meter | flg_w-bath | w-bath | flg_tmp_conversion | tmp_conversion | std_ph4_ph | std_ph4_chem | std_ph7_ph | std_ph7_chem | water_tmp_cal | de-ionized water_ec
pH | Vu Xuan Hung | 5 | 5 | 2021-12-02 | 2021-12-02 | 6.25 | 75 | 25 | 5.05 | 5.05 | 5.04 | 5.0467 | 5.06 | 5.04 | 5.06 | 5.0533 | 5.06 | 5.05 | 5.06 | 5.0567 | 5.0522 | 75 | 25 | 5.48 | 5.49 | 5.5 | 5.49 | 5.5 | 5.47 | 5.49 | 5.4867 | 5.47 | 5.5 | 5.48 | 5.4833 | 5.4867 | Metrohm | 781 | 18 | Metrohm | Metrohm 6.0.258.010 | 9 | 2 | Digital | 1 | Desk top | 1 | Used | 1 | Yes | 4.01 | 6.98 | Merck | Merck | 25 | 0.01
| name | exp1 | exp2 | start_date | finish_date | EC of blank sample | h_serial_no | h_water_tmp | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_water_tmp | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | manfacturer_ECmeter | model_ECmeter | years_ECmeter | manufacturer_cell | model_cell | years_cell | flg_disp | type_disp | flg_meter | type_meter | flg_w-bath | w-bath | flg_tmp_conversion | tmp_conversion | org_std_chem | org_std_conc | dilute1_conc | dilute1_ec | dilute2_conc | dilute2_ec | dilute3_conc | dilute3_ec | water_tmp_cal | cell_const
EC | Vu Xuan Hung | 5 | 5 | 2021-12-02 | 2021-12-02 | 0.01 | 75 | 25 | 2.93 | 2.9 | 2.92 | 2.9167 | 2.91 | 2.89 | 2.92 | 2.9067 | 2.93 | 2.88 | 2.92 | 2.91 | 2.9111 | 75 | 25 | 1.21 | 1.17 | 1.18 | 1.1867 | 1.22 | 1.18 | 1.2 | 1.2 | 1.2 | 1.19 | 1.19 | 1.1933 | 1.1933 | YSI | 3200 | 18 | YSI | YSI 3252 | 18 | 1 | Analogue | 1 | Desk top | 1 | Used | 1 | Yes | HANNA |  |  | 8.4 |  |  |  |  | 25 | 1
| name | exp1 | exp2 | start_date | finish_date | SO42- conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | heat_tmp | heat_time | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
SO42- | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.028853 | 75 | 49.8 | 49.4 | 49.7 | 49.6333 | 50.1 | 48.9 | 49.6 | 49.5333 | 50.5 | 48.4 | 49.2 | 49.3667 | 49.5111 | 75 | 20.2 | 20 | 19.7 | 19.9667 | 19.8 | 20.1 | 19.9 | 19.9333 | 20.2 | 19.6 | 19.5 | 19.7667 | 19.8889 | 0.057706 | 0.19235 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep - Anion dual 2 | Metrosep RP2guard/3.5 | NaHCO3 &Na2CO3 | 0.8 | 2 | Manual | 20 | 35 | 1 | #NAME? | 1 | Attached | 753 Suppressor Module - Metrohm |  |  |  | Na2SO4 | 1000 | 1 | mg/L | 1 | Commercial | 2018 | 2019 | 2000
| name | exp1 | exp2 | start_date | finish_date | NO3- conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | heat_tmp | heat_time | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
NO3- | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.038243 | 75 | 39.3 | 38.9 | 39 | 39.0667 | 38.6 | 38.5 | 38.9 | 38.6667 | 38.7 | 38.4 | 39.2 | 38.7667 | 38.8333 | 75 | 13.4 | 12.9 | 13.4 | 13.2333 | 13.3 | 13.1 | 13.5 | 13.3 | 13.1 | 13.4 | 13.5 | 13.3333 | 13.2889 | 0.076485 | 0.25495 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep - Anion dual 2 | Metrosep RP2guard/3.5 | NaHCO3 &Na2CO3 | 0.8 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 1 | Attached | 753 Suppressor Module - Metrohm |  |  |  | NaNO3 | 1000 | 1 | mg/L | 1 | Commercial | 2018 | 2019 | 2001
| name | exp1 | exp2 | start_date | finish_date | Cl- conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | heat_tmp | heat_time | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
Cl- | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.033541 | 75 | 43.5 | 43.1 | 43.4 | 43.3333 | 42.9 | 43.6 | 43.7 | 43.4 | 43 | 43.4 | 43.1 | 43.1667 | 43.3 | 75 | 19.2 | 18.4 | 19 | 18.8667 | 18.5 | 18.8 | 18.2 | 18.5 | 18.1 | 18.3 | 18.4 | 18.2667 | 18.5444 | 0.067082 | 0.22361 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep - Anion dual 2 | Metrosep RP2guard/3.5 | NaHCO3 &Na2CO3 | 0.8 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 1 | Attached | 53 Suppressor Module - Metrohm |  |  |  | NaCl | 1000 | 1 | mg/L | 1 | Commercial | 2018 | 2019 | 2001
| name | exp1 | exp2 | start_date | finish_date | NH4+ conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | heat_tmp | heat_time | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
NH4+ | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.038827 | 75 | 55.6 | 55.1 | 55.4 | 55.3667 | 54.9 | 55.4 | 54.6 | 54.9667 | 55.8 | 55.5 | 55.2 | 55.5 | 55.2778 | 75 | 23.7 | 23.4 | 23.8 | 23.6333 | 22.9 | 23.5 | 23.7 | 23.3667 | 23.1 | 23.4 | 23.7 | 23.4 | 23.4667 | 0.077653 | 0.25884 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep cation 1-2 | Metrosep C6 S- guard/2.0 | Acid Tactric & Acid Dipinilic | 1 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 0 | Not attached |  |  |  |  | NH4Cl | 1000 | 1 | mg/L | 1 | Commercial | 2019 | 2019
| name | exp1 | exp2 | start_date | finish_date | Na+ conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | flg_correction | correction_method | flg_cesium | cesium_solution | flg_flame | flame | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
Na+ | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.0502 | 75 | 33.9 | 34.5 | 34.3 | 34.2333 | 33.7 | 34.8 | 34.1 | 34.2 | 33.6 | 35.1 | 34.5 | 34.4 | 34.2778 | 75 | 14.3 | 13.9 | 14.1 | 14.1 | 13.8 | 14.5 | 14.3 | 14.2 | 14.1 | 14.4 | 13.9 | 14.1333 | 14.1444 | 0.1004 | 0.33466 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep cation 1-2 | Metrosep C6 S- guard/2.0 | Acid Tactric & Acid Dipinilic | 1 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 0 | Not attached | [8 empty] | NaCl | 1000 | 1 | mg/L | 1 | Commercial | 2019 | 2019
| name | exp1 | exp2 | start_date | finish_date | K+ conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | flg_correction | correction_method | flg_cesium | cesium_solution | flg_flame | flame | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
K+ | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.031105 | 75 | 6.8 | 6.5 | 6.3 | 6.5333 | 7.1 | 6.2 | 6.4 | 6.5667 | 6.5 | 6.3 | 6.4 | 6.4 | 6.5 | 75 | 3.9 | 4.2 | 3.8 | 3.9667 | 4.1 | 4.4 | 3.7 | 4.0667 | 3.8 | 3.7 | 4.6 | 4.0333 | 4.0222 | 0.062209 | 0.20736 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep cation 1-2 | Metrosep C6 S- guard/2.0 | Acid Tactric & Acid Dipinilic | 1 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 0 | Not attached | [8 empty] | KCl | 1000 | 1 | mg/L | 1 | Commercial | 2019 | 2019
| name | exp1 | exp2 | start_date | finish_date | Ca2+ conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | flg_correction | correction_method | flg_lanthanum | lanthanum_solution | flg_flame | flame | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
Ca2+ | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.069875 | 75 | 23.8 | 23.3 | 23.6 | 23.5667 | 24.1 | 23.4 | 22.9 | 23.4667 | 23.6 | 23.5 | 23.4 | 23.5 | 23.5111 | 75 | 8.5 | 8.9 | 8.6 | 8.6667 | 8.4 | 8.2 | 9 | 8.5333 | 9.1 | 9.4 | 8.8 | 9.1 | 8.7667 | 0.13975 | 0.46583 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep cation 1-2 | Metrosep C6 S- guard/2.0 | Acid Tactric & Acid Dipinilic | 1 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 0 | Not attached | [8 empty] | CaCl2 | 1000 | 1 | mg/L | 1 | Commercial | 2019 | 2019
| name | exp1 | exp2 | start_date | finish_date | Mg2+ conc of blank sample | h_serial_no | h_a_data1 | h_a_data2 | h_a_data3 | h_a_average | h_b_data1 | h_b_data2 | h_b_data3 | h_b_average | h_c_data1 | h_c_data2 | h_c_data3 | h_c_average | h_total_average | l_serial_no | l_d_data1 | l_d_data2 | l_d_data3 | l_d_average | l_e_data1 | l_e_data2 | l_e_data3 | l_e_average | l_f_data1 | l_f_data2 | l_f_data3 | l_f_average | l_total_average | nd | ldl | Std 0 | Std 1 | Std 2 | Std 3 | Std 4 | Std 5 | Std 6 | Std 7 | Std 8 | Std 9 | Std 10 | analytical_cd | analytical_method | manufacturer_apparatus | model_apparatus | years_apparatus | analytical_column | guard_column | eluent | flow_rate | flg_injection | injection_method | injection_vol | column_tmp | flg_detector | detector | flg_suppressor | suppressor | model_suppressor | wave_length | flg_correction | correction_method | flg_lanthanum | lanthanum_solution | flg_flame | flame | std_chem | std_conc | flg_unit | unit | flg_preparation | prepatation | ana_col_inst | gur_col_inst | supp_inst
Mg2+ | Nguyen Truong Giang | 8 | 8 | 2021-12-02 | 2021-12-02 | 0.057511 | 75 | 11.3 | 11.1 | 10.9 | 11.1 | 10.5 | 10.9 | 11 | 10.8 | 10.9 | 10.7 | 11 | 10.8667 | 10.9222 | 75 | 4.3 | 4 | 3.7 | 4 | 3.8 | 4.2 | 3.6 | 3.8667 | 3.9 | 4.2 | 3.8 | 3.9667 | 3.9444 | 0.11502 | 0.38341 | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? | 1 | Ion Chromatography | Metrohm | Mic 3 | 1998 | Metrosep cation 1-2 | Metrosep C6 S- guard/2.0 | Acid Tactric & Acid Dipinilic | 1 | 2 | Manual | 20 | 35 | 1 | Electric Conductivity | 0 | Not attached | [8 empty] | MgCl2 | 1000 | 1 | mg/L | 1 | Commercial | 2019 | 2019
| Q1 | Q2_1 | Q2_2 | Q2_3 | Q2_3_o | Q3_1_a | Q3_1_b | Q3_1_c | Q3_1_d | Q4_1_a | Q4_1_a_o | Q4_1_b | Q4_1_b_o | Q4_1_c | Q4_1_c_o | Q4_2_a | Q4_2_a_o | Q4_2_b | Q4_2_b_o | Q4_2_c | Q4_2_c_o | Q4_2_d | Q4_3_a | Q4_3_a_o | Q4_3_b | Q4_3_b_o | Q4_3_c | Q4_3_c_o | Q4_3_d | Q4_4_a | Q4_4_a_o | Q4_4_b | Q4_4_b_o | Q4_4_c | Q4_4_c_o | Q4_4_d | Q5
Questionnaire | 1 | 1 | 1 | 0